{"text_format":"2","workbook":"xlsx","n_sheets":3,"sheets":["Sheet1","Sheet2","Sheet3"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":23,"n_cols":2,"rows_kept":23,"cut":false,"hidden":false,"merged":[]}
["序号","采购要求的技术条款"]
[1,"仪器名称：血气生化分析仪；"]
[2,"方法学：微流控干式电化学法；"]
[3,"电极测量方式：免维护微电极技术；"]
[4,"进样方式：自动液压式进样；"]
[5,"测试参数：pH 、pCO2 、pO2、K+ 、Na+ 、Cl- 、iCa2+、Glu 、Lac 、Hct（一张测试卡可同时检测至少10项实测参数）；"]
[6,"计算参数：cH+(T)、pH(T)、pCO2(T)、pO2(T)、pO2(A-a)(T)、pO2(a/A)(T)、RI(T)、pO2(T)/FiO2、cH+、iCa2+（7.4）、HCO3-act、BE(ecf)、BE(B)、BB(B)、AGsO2(est)、tHb(est)、TCO2、pO2(A-a)、pO2(a/A)、RI、mOsm、pO2/FiO2；"]
[7,"标本类型：可适用于动脉血、静脉血等；"]
[8,"定标方式：测试卡液包定标，无需额外试剂包，无需额外定标液；方便临床操作，血气生化测试卡自带定标液，一体化耗材。"]
[9,"测试卡：血气生化测试卡常温质保≥8个月"]
[10,"质控体系：电子质控卡和质控品两种质控方式，提供原厂配套三级质控品。"]
[11,"操作界面≥8寸电容彩色触摸屏操作，中、英文语言自由切换，内置操作教程；"]
[12,"设备重量：小巧便携≤3.5 kg（含电池）；"]
[13,"内置电池：可充电锂电池，连续待机24小时，连续检测次数≥50次；"]
[14,"仪器配置：仪器内置条码扫描仪及热敏打印机；"]
[15,"数据接口：串口、网络接口、有线、无线网络链接，可直接连接LIS、HIS系统；数据管理：仪器可自动存储100000个病人结果，10000个质控结果，连接POCT数据管理系统。"]
[16,"参数升级：院内联网升级软件，使用新的测试卡即可完成，无需增加模块；"]
[17,"实际用血量：实际用血量≤150 uL；"]
[18,"测量时间：样本检测时间不超过60s。"]
[]
[null,"采购要求的商务条款"]
[1,"质保要求：自开始使用（装机）之日起≥3年内免费保修，电池保修6个月."]
[2,"供应商接到保修请求，维修应在1小时内响应，4小时内维修人员到达现场，如遇重大问题或者其他暂时无法迅速解决问题在5个工作日内解决，必要时应向采购人提供应急备用设备。"]

{"sheet":"Sheet2","index":1,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[]}

{"sheet":"Sheet3","index":2,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[]}